## Overview
1 | Document name | Request for quotation
2 | Related document
3 | Revision History
|  | Version | Content
|  | 0.1 | Initial version
|  | 0.2 | Update sales country
4 | Introduction
Purpose of use or effect
Scope of application
Glossary
References
5 | Product Overview
Overview
Main function
Product composition
Sales country
USA
EU countries
Legal regulations
ISO13485（Medical devices - Quality management systems - Requirements for regulatory purposes）
ISO14971（ the application of risk management to medical devices）
IEC60601-1 ed4 (General requirements for basic safety and essential performance)
IEC62366（Application of usability engineering to medical devices）
IEC62304 (software Life Cycle Processes)
Compliance with CE、FCC、UL
IEC62133（Secondary cells and batteries）

## Detail Requirement
1 | Market/Environment
|  | SAM(Served Available Market) | USA
EU countries
Operational Environment condition
Ambient temperature: 0℃~+40℃
Relative humidity:Below 85%RH(No dew condensation)
IP(International Protection) Code : IP67 or greater
Classification by protection type against electric shock (IEC 60601-1)
Power cable use : Class １
|  | Storage Environment condition | Ambient temperature: 0℃~+55℃
Relative humidity:Below 85%RH(No dew condensation)
|  | Storage Period | > 6 years
2 | Hardware spec | External dimensions | W150×H70×D90mm
|  | weight | Body : less than 700g | 努力するが設計完了時質量は超える恐れがある
|  | Display | 4.3inch or more, 16:9 | 4.3inch 16:9 95.1 x 53.9
|  |  | Touch panel | ユーザビリティ規格を参照
Number of colors:256 or more
Viewing angle:Left right130°or wider/Updown130°or wider
|  | switch | Power switch is required
Power switch is in a location that can not be pressed during operation
Run/stop button is required
|  | Color | Black
|  | Material | ABS resin
Flame retardant grade V0 expected
Resistance for yellowing
Use antibacterial resin
|  | Power supply | AC100V～240V 50Hz/60Hz 3pins
Internal AC/DC power supply
Use power cable and connector that meet the standards of medical equipment
Connector shape：IEC-60320-C13/14
|  | Rechargeable battery | Continuous run time > 5 days
Full charge time < 6 hours
Rechargable batteriy can be charged while power cable is connected.
Rechargable battey shall be available for air trasportation.
IEC62133 is complied.
Battery shall be able to replaceable
Sub power supply
Continuous standby time : > 15 minutes min.
It is installed in equipment.
It is a power supply to prevent power outage while replacing the battery.
When AC power supply and battery are lost, power lost alarm warns.
Interface
USB port ×1
LAN Connector×１, RJ-45 LAN port（10/100Mbps or more）
|  | Wifi | IEEE802.11b/g/n/a/ac (2.4GHz/5GHz) shall be satisfied | AP6255
|  | Internal memory | >20 days measurement results
|  | Speaker | Emergency speaker shall sounds when abnormal circumstances
The volume can be changed in 3 steps
|  | Accelerometer | Shall be able to detect the tilt of the equipment by the acceleration sensor
Inclination accuracy : 1°
3 | Software spec | Non measurement | When turning on the power switch, the power turns on and becomes the standby screen
Shall be able to display date and time.
Shall be able to display icon that radio field strength.
Shall be able to display icon that remaining capacity of the battery
Shall be able to display icon that connecting AC
* Detailed specifications are defined by software specifications
|  | Parameter setting | Shall be able to set Wi-Fi connection
(SSID, password, security authentication, IP address, subnet mask,
Gateway IP address, MAC address display, 2.4 GHz / 5 GHz selection)
Shall be able to turn Wi-Fi ON / OFF
Shall be able to adjust date and time using Wi-Fi
Shall be able to adjust the date and time
Shall be able to turn identification ON / OFF
Shall be able to set display no date and time
Shall be able to initialize
* Detailed specifications are defined by software specifications
|  | error | Shall be able to display error contents when error happen
Shall be able to send error contents using Wi-Fi (only Wi-Fi ON)
Notify by buzzer
The details of the error are as follows
· There is no power supply from the AC adapter and the battery is low
· There is no power supply from the AC adapter and press "RUN" button
· Misplacement of battery
4 | Other | Drop Impact resistance | Shall be no functional loss by dropping from 1 m height to the flat concrete floor
Bottom: 1 side 4 sides 4 corners
|  | packing | One box per each
|  | Environmental Progress | RoHS compliant
|  | Other | It is possible to update FW through Wi-Fi
5 | Accessories | battery recharge tarminal | Terminal that can charge the main battery
|  |  | The ability to charge more than 4 terminals with one terminal | Eneroid
In AC100V～240V 50Hz/60Hz
|  | Mounting bracket | Mounting bracket that can be attached to drip pole
|  | Power cable | 3 pin connector
2m and longer

## Quotation Requirement
1 | NRE ( non-recurring expense )
Hardware development
ID/ME development
Hard Tooling
Test fixture for mass production
2 | Unit price for the following quantity per order
1K
5K
20K
3 | Proposed solution block diagram / key component
4 | Milestone and schedule